Relative Efficiency (at 1.33 MeV)

For historical reasons, the relative detection efficiency of coaxial germanium  detectors is defined at 1.33 MeV relative to that of a standard 3-in.-diameter, 3-in.-long NaI(Tl) scintillator. The measurement is performed by the method that is described in the IEEE Standard Test Procedures for Germanium Detectors for Ionizing Radiation (ANSI/IEEE 325–1996) and in the equivalent IEC standard. 

A National Institute of Standards 60 Co source with known intensity is positioned 25 cm from the endcap face, and a fixed-time count is taken for the 1.33-MeV peak. The absolute efficiency is the ratio of the number of counts in the photopeak divided by the number of gamma rays emitted from the source during the same period of time. This absolute efficiency is then divided by 1.2 X 10–3 , which is the absolute efficiency at 1.33 MeV of a standard 3-in. by 3-in. NaI(TI) crystal 25 cm from the source. 

The ratio of these measurements is the basis for the relative efficiency specification of the germanium detector. Relative efficiency, while giving a general indicator of detector performance, can be highly misleading in regards to specific geometries (e.g., filter paper or Marinelli beakers). 

[image: image1.jpg]



Figura 1 ORTEC GEM-40210. Note that 40 is the relative efficiency at 1.33 MeV= 40 % to a NaI (Tl) 3"x3" and 210 is the Energy resolution at 1.33 MeV i.e. FWHM = 2.10 keV
Impulsi e affini vedi:

http://it.emcelettronica.com/tecniche-di-condizionamento-del-segnale-spettrometria-gamma

