	MISURA DELLA CONCENTRAZIONE DI RADIOATTIVITA’

Rapporto di prova n° 1.2

	Data di rilascio:  4 luglio 2016

	Committente: 

	Descrizione campione: pezzi variegati di ceramiche

	Prelievo del campione:

Prelievo:    Data indefinita                                      Modalità di prelievo:   A cura del committente

Consegna: 20 maggio 2016                                    Modalità di consegna: A mano



	Strumentazione:

Rivelatore:             HPGe ORTEC GEM-08180-S Risoluzione 1.72 (FWHM) a 1.33 MeV, Co-60

Catena elettronica: NIM-BIN ORTEC: 401A (crate); 459 (HV); 572 (amp.) MCA TRUMP 8k; Software:               GammaVision V6.8

Schermatura:          Pozzetto in piombo spessore 5 cm rivestito all’interno con 0,5 mm rame

Calibratore:            Radionuclide: Becher Marinelli con Eu-152 attività 3050 Bq al 15/06/2016 

                               tipo MBSS 1 SN: 190516-1529003 Czech Metrology Institute (vedi allegato)



	Norme di riferimento:

UNI 10797:1999 “Radionuclidi naturali nei materiali da costruzione – Determinazione mediante spettrometria gamma ad alta risoluzione”



	Risultati:  Data misura:  1 luglio 2016   Durata della misura: Real time: 1916 s; Live time: 1914 s



	Radionuclide
	Energia

(keV)
	Concentrazione

(Bq/kg)
	Errore statistico

(Bqkg)
	MDA

(Bqkg)

	Ra-226
	186,1
	157,3
	53,3
	103,8

	Pb-214
	242,0
	58,6
	14,6
	37,8

	Pb-214
	295,2
	61,5
	12,5
	15,9

	Pb-214
	351,9
	62,0
	7,1
	9,4

	Bi-214
	609,3
	60,5
	6,8
	14,5

	Bi-214
	1764,49
	77,2
	27,0
	34,6

	Pb-212
	238,6
	62,7
	5,0
	16,5

	Tl-208
	510,8
	33,5
	10,7
	10,9

	Tl-208
	583,1
	19,1
	2,9
	7,1

	Ac-228
	911,2
	76,2
	17,6
	26,8

	K-40
	1460,8
	785,3
	70,1
	80,7

	Attività totale da 10,3 a 1982,4 keV = (913,4 (87,5) Bq/kg

	Indice di radioattività – Radiation Protection 112 della Commissione Europea (ISBN 92-828-8376-0)

             Valore misurato                                                        Valori di riferimento

                                 I = 0,62 ( 0,08                                                    I ( 0,5 esenzione     I (1,0 controllo




                                                                                                                          Prof. Valter Giuliani

Allegato 1

[image: image1.jpg]Counts.

20
20
20
20
210
2
130
180
170
180
150
140
130
120
110
100
@
a
n
&
E
a
E
x
10

SPETTRO DI EMISSIONE GAMMA DEL CAMPIONE DI PEZZI DI CERAMICHE ( Real/Live Tim:

918/1913 (5).)

Pb-212

Ac-228

Bi-214]

oy

100

200

00

an0

00

600

0

a0 w0 tom
Energy (keV)

1100

1200

1300

1400

1500

180 7m0 1s00




Spettro di emissione gamma

[image: image2.jpg]Efficienza

0,02500

0,02000

001500

001000

0,00500

0,00000

Becher Marinelli con Eu-152

nza

— it polinomiale 3° ordin
= valori sperimentali

200

400

600

800
Energia (keV)

1000

1200 1400

1600




Allegato 2

[image: image3.jpg]¢ oby- 796
V. Czech Metrology Institute

Okruzni 31, 638 00 Brno, Czech Republic
phone +420 545 555 111

www.cmi.cz

Workplace: Regional Branch Prague, Radiova 1136/3, 102 00 Praha 10
Ionizing Radiation Building, Radiova 1288/1a
Phone: +420 266 020 497

CERTIFICATE

Certificate No.: 1035 - SE - 40444-16  Type: MBSS 1 Serial No.: 190516-1529003

Half life Activity Combined standard
Radionuclide days kBq uncertainty, %
Eu-152 49388 3,050 13

Mass: 4410 g Density: 0,98 + 0,01g/cm3 Volume: 450,0 + 45 cm3
Radionuclide impurities: Eu-155 0,02 %

Reference date: 15.6.2016 Homogeneity better than: 1%
Description:

Radioactive material is homogeneously dispersed in silicone resin. Composition of the
matrix: C-0,324 H-0,0816 O-0,216 Si- 0,379 (mass ratio).

Measuring method:

Preparation issues from standard ER solutions whose activities were determined by
suitable absolute method. Final control is based on gamma spectrometry on HPGe
detector.

Note:

As the criterion of homogeneity standard deviation of the activity value of 1 cm”3
element was chosen (n=10). The volume is calculated from the mass and the density.

Date of the certificate issue: 6.6.2016

Certificate validity: 3 yeaxy

e




ORTEC GEM- 08180- S DWR-1.2G (1,2 litre)

General description

The ORTEC GEM-08180-S is a P-type coaxial HPGe detector with 30 liter dewar.

The majority of gamma spectroscopy applications such as those found in counting laboratories involve the energy range ~40 keV upward. ORTEC GEM and is designed to cover this energy range.

These detectors are manufactured from ORTEC-grown germanium crystals processed in their advanced manufacturing facility in Oak Ridge, TN. The detectors are fabricated from P-type germanium with an outer contact of diffused Li and an inner contact of ion-implanted boron.

Specifications

	Relative efficiency at 1332 keV to a NaI(Tl) 3”x3”
	 8%

	Energy resolution at 1332 keV
	1,83 keV


[image: image4.jpg]Log (Counts)

st
Channels





