DETECTION EFFICIENCY

[image: image1.jpg]Full-energy Peak Efficiency (FEPE): &(E)

o of the number of counts in full-energy peak corresponding
rgy E (Np(E)), by the numberoﬁphotons nergy £

Counts per
channel

Energy

£-(E) depends on the source-detector geometry
and on the energy




.............

[image: image2.jpg]Full-energy Peak Efficiency (FEPE): &5(E)

€,(E) depends on the geometrical conditions and on the energy

Detecto r

characteristics
-QM Intrinsic
efficiency

Shielding





..............

[image: image3.jpg]Associated standard uncertainty

Efficiency calibration £(E) = N(E)

Influence of the peak area :

if N=10* AN/N=102-> Ag/e=1,1102

if N=105 AN/N=3,110%-> Ae/e=6103




..........

[image: image4.jpg]Geometrical efficiency

d
Detector Q=2 7;'(1 - d )
d? +r?

emitted towards the detector
emited by the source

)
¢ 4

geomelry




..............

Note

1) I((E)= The branching ratio (or yield) is used to correct the number of gamma rays observed to obtain the number of disintegrations of the source

2) in cases when the standard geometry is identical to the counting geometry,  the correction due to geometrical differences becomes virtually negligible.

3) The corrected net peak area (N) for the background was obtained according to the expression:


[image: image5.wmf],

,

Ltotal

totBG

LBG

t

NNN

t

=-


 4) In the case of materials characterized by high radon exhalation, the sealing is very important in order to reduce the 222Rn loss. The sealing effectiveness and consequently the 222Rn growth within the container were successfully checked, as shown in Fig. where the in-grow of count rates of radon progeny 214Bi (at 609 keV) is displayed for a phosphogypsum sample. The in-growth counts were measured for six 222Rn half-lives (corresponding to approximately 99 % of equilibrium). The experimental data points are well fitted (with a reduced (2 = 1.0) taking the 222Rn half-live (3.821 days0 as a fixed parameter.

[image: image6.jpg]214gj (609 keV) counts

40000

39000

g
8

37000

3
222Rn half-lives





5) The energy and FWHM (Full Width at Half Maximum) determined for the most intense photopeaks are well fitted with a first order (Eq. 1) and a second order (Eq. 2) polynomial function, respectively, with a reduced v2 = 1.0. Eð Þ¼ keV a1 Channel þ 


[image: image7.wmf]2

012

2

012

()

()

EkeVaaChannelaChannel

FWHMkeVaaChannelaChannel

=+×+×

=+×+×


for example, in the Gammavision report:

[image: image8.jpg]Energy calibration

Created: 20/03/2017 16:31:11
Zero offset: 0.369 kev
Gain: 0.243 kev/channel

Quadratic: -3.0796-08 kev/channelA2




For the FWHM:

[image: image9.jpg]FWHM{[keV]=4.4484-+0.000782x(Channel)-1.50488e-008(Channel)?, and it
is shown in Figure 3, [6, 8].

Fig. 3 — The calibration in resolution curve of the system.




6) The final sensitivity of the measurements can be evaluated by using the detection limit (LD), assuming the Gaussian probability distribution of the number of counts in the background (B) and rejecting the data not included in a range of 1.645r (95 % confidence level): 


[image: image10.wmf]2.714.654.65

D

LBB

=+

;


where the approximation is admitted for high number of counts. The minimum detectable activity (MDA) for the background is calculated using the LD, according to the formula:


[image: image11.wmf]D

L

MDA

It

g

e

=


FITTING EFFICIENCY

P-type germanium detectors, such as the ORTEC GEM Series, have a maximum efficiency at about 150 keV; for n-type detectors, such as the GMX Series, it is about 100 keV. For detectors above about 50% relative efficiency, these values will be somewhat higher (Fig. 107). For both types, these maxima, or knee values, depend on the individual detector. For p-type GEM detectors, the efficiency goes down as the energy goes down from the knee. For n-type GMX detectors, the efficiency is nearly constant at energies below the knee. For both types, the efficiency goes down at energies above the knee.

[image: image12.jpg]Logarithmic Efficiency (Arbitrary Units)

Offset for Purposes of Comparison

THIN WINDOW

000
Energy (Log)

Detector Efficiency as a Function of Energy




.........

[image: image13.jpg]w

. Interpolative fit.

Linear fit of the natural logarithm of the efficiency to the natural LnlEfpola‘iVB
1 F ; Inear
logarithm of the energy. Guadratic
Palynomial )
. Quadratic fit of the natural logarithm of the efficiency to the natural Tcc Polynomial =

logarithm of the energy. Fig. 108.

Polynomial fit of the natural logarithm of the efficiency to the energy."®

. TCC Polynomial, a different six-order polynomial fit of the natural logarithm of the

efficiency to the natural logarithm of the energy




Options 1, 2, and 3 can be selected separately for two separate energy regions. Either of the two regions might be left uncalibrated by not including any points in the region, but the analysis will report zero intensity (in the library peak output) for peaks in the uncalibrated region. If both regions are calibrated, the above-the-knee energy region is fitted first, and the calculated efficiency at the knee is included as a data point in the below-the-knee fit. This means that only one point need be below the knee, but two points are the minimum above the knee for a calibration to be done. Option 4 fits the entire energy range with one function and is best suited to p-type detectors. Option 5 fits the entire energy range with different functions over three energy regions and can be used for p- or n-type detectors.

[image: image14.jpg]Linear Fit

The linear fit uses a straight-line fit to the data points. This 1s used when few data points are
used or if the data points are all very close in energy. Also, for an n-type detector, the below-the-
knee efficiency is nearly a straight-line down to approximately 10 keV., and the linear fit will
produce the best result. For only two data points in the entire calibration, linear and interpolative
are equivalent. If sufficient data are available, the linear fit is not as accurate as the quadratic fit

because of the efficiency knee.
The efficiency/energy formula 1s:

= e<”1’(’2‘h‘(E))




.......

[image: image15.jpg]Quadratic Fit

The quadratic fit fits a quadratic function to the log (energy) vs. log (efficiency) curve. At least
three data points above the knee and two below the knee are required for this fit. With only three
points, the fit will be reported as exact for all data points, but the calibration might be inaccurate
elsewhere.

If the input points are not well separated, the best fit to the data points might not be an accurate
representation of the efficiency outside the fitted region.

The efficiency/energy formula is:

&t (ay=ay ln(E)fa}(ln(E))z)

where:
¢ = efficiency at energy E
a; = fitting coefficients

E = energy




........

[image: image16.jpg]Efficiency calibration

Created: 07/05/2015 09:21:16

Knee Energy: 20.00 kev

Above the Knee: qQuadratic Uncertainty = 0.52 %

Log(eff): -8.5632736-01 + ( 1.465502E-017Log(E) ) +
(27.236815€-02%L0g ()2 )

Belon the knee: Linear Uncertainty =  0.00 %

Log(eff): 0.000000E+00 + ( 0.000000E+00*Log(E) ) +

¢ 0.000000E+00%Log ()42 )




.........

Polinomial fit

[image: image17.jpg]grafico LnE_Lneff

y = 0.1149x* - 2.2423x" + 13.546x - 29.823

Lneff

+ Seriel

—Poli_(Seriet

LnE





.........

[image: image18.jpg]) Microsoft Excel - Efficienza D AT 10716 tris

Fie Modfica Visualzza Inserisci Formato Strumenti Dafi Festra 2

DEHESRY tBRS v- @ = £ 4| Mm B 100 - @) .| Al S0 -|GC
09 - =| =EXP(SJSATLN(9F3+SKS4"LN(9P2+SLEALN(9)+SNIS4)
A 5 © [) E F G ] ]

1 E[KeV] Effcienza LnE Lneft

2

3| 12176 232E-02 4802051894 3763603 0.6 34586 202 0000866418 %3 %2
4 24471 178E-02 6500073836 40303 0.076 7282 102 0.000768047 01149

5| 34428 13302 584145528 4321647 0.7 18853 145 0.000479383

6| 41111 114E02 6018860819 4478538 0022 1231 54 0.000752856

7| 44386 106E-02 6.005609197 4547279 0031 1588 57 0.000614842

8| 6883 685E-03 6534224789 4983857 001  313.00  44.00 0.000972495 E off Polinom_eff
9| 77887 56503 6657844151 6175604  0.13 3936 7 0.000271935 20__o00788] ¢
10 867 530E.03 6765038977 5240048  0.042 1139 54 0.000382384 30000569
11| 96391 4.98E-03 6670997929 -6.303028 015 3741 2 0.000223716 40 001007
12| 100508 4.57E-03 6912772671 -6.388242 0006 179 3 0001225662 50 001401
13| 108591 4.24E-03 6.990173624 -6.463192 01 21 61 0.00024286 5 001570
14| 11121 4.09E-03 7.014005399 -6498697 014 3056.00  65.00 0.000203851 60 001718
16| 121295 381E-03 7.100810688 -6.570126 001  260.00  30.00 0.000491803 70 001956
16| 12991 3.58E-03 7.1694269% 6632392 002  294.00  25.00 0000385018 80 002122
17| 1408.04) 3.33E03 7.249953945 -6703613 021 3765.00  63.00 0.000152615 80 002122

(

18

100

0.02297




............

GammaVision Polinomial fit 

[image: image19.jpg]Polynomial Fit

The polynomial fit uses a 6-term polynomial to fit the natural logarithm of efficiency to the
energy. The function is optimized for p-type detectors. For n-type detectors. the low-energy
region (below 60 keV) is not well modeled by this function. At least five well-separated peaks
are needed for the polynomial fit. The knee value is not used in this fit and the function is
grayed. The polynomial option is only on the Above knee list: the Below knee option is not used
if polynomial is selected.

The polynomial efficiency/energy formula is:

where:
¢ = efficiency at energy E
a, fitting coefficients
E = energy in MeV

The result of the efficiency calibration calculation is one or two sets of coefficients (one for the
fit above the knee and one for below the knee, or just one for the polynomial fit). and a set of
energy-efficiency pairs. The energy-efficiency pairs are used for the interpolative fit. The pairs
might also be used to recalculate the efficiency and to display the efficiency plot.




...


[image: image20.wmf]1234

123456

ln

aEaaEaEaEaE

ggggg

e

----

=+++++


energy in MeV !

..

[image: image21.jpg]DEHESRY 4 BBRI o-c-|@= A 45 [0S 10% -0 cw-lges)

Ke - =EXP(0.131749°M9-6.390524+ 1 477391"M9"1-0.385984"M"-2+0.052091"NI9"-3-0.002706'M"4)
— A | B c = E H || g I K T
E [KeV] Efficienza LnE Ln eff I N EmN EnEff
121.76. 2.32E-02 4.802051894 -3.763603 0.26 34586 202 0.000868418 x*3 x*2 x
24471 1.78E-02 5500073836 4.0303 0.076, 7282 102 0.000768047 0.1149 -2.2423 13.5¢
34428 133E-02 584145528 -4.321647 027 18853 145 0.000479383
41111 1.14E-02 6.018860819 4478538 0.022 1231 54 0.000752856
44386 1.06E-02 6.095509197 -4.547279 0.031 1588 57| 0.000614842
688.3 6.85E-03 6.534224789 -4.983857 0.01 313.00 44.00 0.000972495 E eff Polinom  eff GammaVis
778.87 5.65E-03 6.657844151 -5.175604 013 3936, 76/ 0.000271935 20 0.00188]_0.000000000%
867 5.30E-03 6.765038977 -5.240048 0.042 1139, 54 0.000382384 30/ 0.00569  0.000000000
963.91 4.98E-03 6.870997929 -5.303028' 0.15, 3741 72 0.000223716 40 0.01007  0.000000000
1005.03 4.57E-03 6.912772671 -5.388242 0.006 179 35 0.001225562 50/ 0.01401  0.000000000
1085.91 4.24E-03 6.990173624 -5.463192 01 2411 61 0.00024286 56 0.01570 0.000000000
11121 4.09E-03 7.014005399 -5.498697 014 3056.00 65.00 0.000203851 60/ 0.01718  0.000000000
1212.95 3.81E-03  7.100810688 -5.570126 0.01 260.00° 30.00 0.000491803 70 0.01955  0.000000000
1299.1 3.58E-03 7.169426996 -5.632392 0.02 294.00 25.00 0.000385018 80/ 0.02122 0.000003482
1408.04 3.33E-03 7249953945 -5.703613 021 3765.00 63.00 0.000152615 80/ 0.02122 0.000003482

100 0.02297  0.008532094

120 002334 0.022288141
70 FE T





When plotting the detector efficiency versus the gamma energy with log-log scale, the values get arranged linearly in a wide energy region. Thus, the relation between detector efficiency and gamma energy can be expressed as:


[image: image22.wmf]lnln

abE

e

=+


[image: image23.jpg]absolute efficiency

1E2

. 83
= £ 3
H 3
& A\
1E-3 o
‘GalL) detector (distanca 17.5 mm)
©  Am241,Ce137, MnS4,Co0
o Ho-teom
1E4 T T
100 1000

ElkeV





Gamma vision

_1553616221.unknown

_1553617083.unknown

_1553630849.unknown

_1553618246.unknown

_1553616327.unknown

_1553614726.unknown

